## Slide 1: Ανόργανη Χημεία (Ε)
Ενότητα 7: Φλογοφωτομετρία
Χριστίνα Φούντζουλα
Τμήμα Ιατρικών Εργαστηρίων
Ανοικτά Ακαδημαϊκά Μαθήματα στο ΤΕΙ Αθήνας
Το περιεχόμενο του μαθήματος διατίθεται με άδεια Creative Commons εκτός και αν αναφέρεται διαφορετικά | Το έργο υλοποιείται στο πλαίσιο του Επιχειρησιακού Προγράμματος «Εκπαίδευση και Δια Βίου Μάθηση» και συγχρηματοδοτείται από την Ευρωπαϊκή Ένωση (Ευρωπαϊκό Κοινωνικό Ταμείο) και από εθνικούς πόρους.

## Slide 2: Flame Photometry Flame atomic emission spectrometry
an atomic emission method
routine detection of metal salts, principally Na, K, Li, Ca, and Ba

## Slide 3: Flame Photometry Flame atomic emission spectrometry
Emission is the process by which a higher energy quantum mechanical state of a particle becomes converted to a lower one through a photon resulting in the production of light.
The frequency of light emitted is a function of the energy of the transition. Since energy must be conserved, the energy difference between the two states equals the energy carried off by the photon.
By heating the sample to a high temperature, the excitations are produced by collisions between the sample atoms. This method is used in flame emission spectroscopy
Each element's emission spectrum is unique.

## Slide 4: Flame Photometry Flame atomic emission spectrometry
How can we distinguish absorption spectrum from emission spectrum?
An absroption spectrum has dark lines where the energy is absorbed and the emmision spectrum has bright lines where the emission energy is.
Why is the emission spectrum of an element a line spectrum and not a continuous spectrum?
It's a line spectrum because of the quantization of energy- meaning you only see energy with levels n=1,2,3.... One would never see the energy level n=2.8 for instance- that would be the case if it were continuous rather than a line spectrum.

## Slide 5: Flame Photometry

## Slide 6: Flame Photometry
Solutions are aspirated into the flame
The hot flame evaporates the solvent,
atomizes the metal, excites a valence electron to an upper state.
Light is emitted at characteristic wavelengths for each metal as the electron returns to the ground state
Optical filters are used to select the emission wavelength monitored for the analyte species

## Slide 7: Flame Photometry
The energy that is needed for the excitation is provided by the temperature of the flame (2000-3000 ºC), produced by the burning of acetylene or natural gas (or propane-butane gas) in the presence of air or oxygen.
By the heat of the flame and the effect of the reducing gas (fuel), molecules and ions of the sample species are decomposed and reduced to give atoms, e.g.: Na+ + e-  Na.
Atoms in the vapour state give line spectra.
(there are no covalent bonds there are not any vibrational sub-levels to cause broadening)

## Slide 8: Flame Photometry
Many different experimental variables affect the intensity of light emitted from the flame.
Fuel and oxidant flow rates and purity
aspiration rates
solution viscosity
concomitants in the samples, etc.
Careful and frequent calibration is necessary for good results

## Slide 9: Flame Photometry

## Slide 10: Flame Photometry
The bright doublet known as the Sodium Dlines at 588.9950 and 589.5924 nm.
These lines are emitted in a transition from the 3p to the 3s levels.

## Slide 11: Flame Photometry

## Slide 12: Flame Photometry
“THE APPLICATION OF FLAME PHOTOMETRY TO SODIUM
AND POTASSIUM DETERMINATIONS IN
BIOLOGICAL FLUIDS”
by RICHARD R. OVERMAN AND A. K. DAVIS, Department of Physiology, University of Tennessee College of Medicine, Memphis
(Received for publication, January 3, 1947)
“THE FLAME PHOTOMETER FOR THE MEASUREMENT OF
SODIUM AND POTASSIUM IN BIOLOGICAL MATERIALS”
by PAULINE M. HALD, Department of Internal Medicine, Yale University School of Medicine, New Haven
(Received for publication, October 30, 1946)

## Slide 13: Flame Photometry
The Determination of Sodium and Potassium in Biological Fluid
Protocol: P05-031A
Introduction
The simple measurement of sodium and potassium in various biological fluids using a flame photometer is described, including dilution ratios, interferences and calibration curves.
Materials Required
Equipment
JENWAY Flame Photometer PFP7
Volumetric Glassware
Pipettes
Reagents
Sodium Standard Solution – 1000mg/l (1000ppm) (Jenway Part Number 025 021)
Potassium Standard Solution – 1000mg/l (1000ppm) (Jenway Part Number 025 023)
Deionised Water

## Slide 14: Flame Photometry
Methods
Calibration using a standard curve
1. Set up the instrument as described in the instruction manual.
2. Prepare the following standard solutions: 0.2, 0.4, 0.6, 0.8 and 1.0mg Na/100ml, from the 1000mg/l sodium standard solution, using deionised water as the diluent.
3. Select the sodium filter and aspirate the 1.0mg Na/100ml standard and adjust sensitivity control to obtain a reading of 100.
4. Aspirate deionised water and adjust the zero control to obtain zero reading.
5. Aspirate 1.0mg Na/100ml standard again, re-adjust to 100.
6. Aspirate deionised water again and re-adjust if necessary to zero.
7. Repeat steps 5 and 6 if necessary to obtain 100 on the standard and zero on deionised water.
8. Aspirate the other standard solutions, note the readings and plot a calibration curve.
9. Repeat the above calibration procedure with potassium filter, using standard solutions of the following concentrations: 0.2, 0.4, 0.6, 0.8 and 1.0mg K/100ml, prepared from the 1000mg/l potassium standard solution using deionised water as the diluent.

## Slide 15: Flame Photometry
SERUM POTASSIUM
Sample preparation
Prepare a 1:50 dilution of haemolysis-free serum in deionised water.
Note: The sodium present in the sample will not affect the potassium determination and no allowance need be made for sodium when preparing the standards.
Reagents
1.0mg K/100ml standard
Blank
Deionised water
Procedure
1. Set up the instrument with the K filter and adjust the flame as described in the instruction manual.
2. Adjust the instrument to obtain 100 on the 1.0mg/100ml standard and zero reading on the blank as described above in the calibration section.
3. Aspirate sample and note the reading.
4. From the calibration curve previously prepared for potassium, read off the concentration equivalent to this reading.
5. To obtain mg K/100ml of serum, multiply this concentration by the dilution factor.
6. Aspirate deionised water to remove all traces of the sample, which might otherwise block the nebuliser.

## Slide 16: Flame Photometry
PLASMA SODIUM
Sample preparation
Prepare a 1:500 dilution of the plasma sample in deionised water.
Note: It is not necessary to remove proteins, also potassium and calcium do not interfere in the quantities present in plasma.
Reagents
1.0mg Na/100ml standard
Blank
Deionised water
Procedure
1. Set up the instrument with Na filter and adjust the flame as described in instruction manual.
2. Adjust the instrument to obtain 100 on the 1.0mg/100ml standard and zero reading on the blank as described above in the calibration section.
3. Aspirate sample and note the reading.
4. From the calibration curve previously prepared for sodium, read off the concentration equivalent to this reading.
5. To obtain mg Na/100ml of plasma, multiply this concentration by the dilution factor.
6. Aspirate deionised water to remove all traces of sample, which might otherwise block the nebuliser.

## Slide 17: Flame Photometry
URINE POTASSIUM
Sample preparation
Prepare a 1:500 dilution of fresh urine in deionised water.
Note: Urine does not contain a sufficient concentration of any substance likely to interfere with potassium determination made with the Flame Photometer and true values may be obtained using fresh urine diluted only with deionised water.
Reagents
1.0mg K/100ml standard
Blank
Deionised water
Procedure
1. Set up the instrument with the K filter and adjust the flame as described in the instruction manual.
2. Adjust the instrument to obtain 100 on the 1.0mg/100ml standard and zero reading on the blank as described above in the calibration section.
3. Aspirate sample and note the reading.
4. From the calibration curve previously prepared for potassium, read off the concentration equivalent to this reading.
5. To obtain mg K/100ml of urine, multiply this concentration by the dilution factor.
6. Aspirate deionised water to remove all traces of the sample, which might otherwise block the nebuliser.

## Slide 18: Flame Photometry
URINE SODIUM
Sample preparation
Prepare a 1:1000 dilution of fresh urine in deionised water.
Reagents
1.0mg K/100ml standard
Blank
Deionised water
Procedure
1. Set up the instrument with Na filter and adjust the flame as described in instruction manual.
2. Adjust the instrument to obtain 100 on the 1.0mg/100ml standard and zero reading on the blank as described above in the calibration section.
3. Aspirate sample and note the reading.
4. From the calibration curve previously prepared for sodium, read off the concentration equivalent to this reading.
5. To obtain mg Na/100ml of urine, multiply this concentration by the dilution factor.
6. Aspirate deionised water to remove all traces of sample, which might otherwise block the nebuliser

## Slide 19: Τέλος Ενότητας

## Slide 20: Σημειώματα

## Slide 21: Σημείωμα Αναφοράς
Copyright Τεχνολογικό Εκπαιδευτικό Ίδρυμα Αθήνας, Χριστίνα Φούντζουλα 2013. Χριστίνα Φούντζουλα. «Ανόργανη Χημεία (Ε). Ενότητα 7: Φλογοφωτομετρία». Έκδοση: 1.0. Αθήνα 2013. Διαθέσιμο από τη δικτυακή διεύθυνση: ocp.teiath.gr.

## Slide 22: Σημείωμα Αδειοδότησης
Το παρόν υλικό διατίθεται με τους όρους της άδειας χρήσης Creative Commons Αναφορά, Μη Εμπορική Χρήση Παρόμοια Διανομή 4.0 [1] ή μεταγενέστερη, Διεθνής Έκδοση. Εξαιρούνται τα αυτοτελή έργα τρίτων π.χ. φωτογραφίες, διαγράμματα κ.λ.π., τα οποία εμπεριέχονται σε αυτό. Οι όροι χρήσης των έργων τρίτων επεξηγούνται στη διαφάνεια «Επεξήγηση όρων χρήσης έργων τρίτων».
Τα έργα για τα οποία έχει ζητηθεί άδεια αναφέρονται στο «Σημείωμα Χρήσης Έργων Τρίτων».
[1] http://creativecommons.org/licenses/by-nc-sa/4.0/
Ως Μη Εμπορική ορίζεται η χρήση:
που δεν περιλαμβάνει άμεσο ή έμμεσο οικονομικό όφελος από την χρήση του έργου, για το διανομέα του έργου και αδειοδόχο
που δεν περιλαμβάνει οικονομική συναλλαγή ως προϋπόθεση για τη χρήση ή πρόσβαση στο έργο
που δεν προσπορίζει στο διανομέα του έργου και αδειοδόχο έμμεσο οικονομικό όφελος (π.χ. διαφημίσεις) από την προβολή του έργου σε διαδικτυακό τόπο
Ο δικαιούχος μπορεί να παρέχει στον αδειοδόχο ξεχωριστή άδεια να χρησιμοποιεί το έργο για εμπορική χρήση, εφόσον αυτό του ζητηθεί.

## Slide 23: Επεξήγηση όρων χρήσης έργων τρίτων
Δεν επιτρέπεται η επαναχρησιμοποίηση του έργου, παρά μόνο εάν ζητηθεί εκ νέου άδεια από το δημιουργό.
©
διαθέσιμο με άδεια CC-BY
διαθέσιμο με άδεια CC-BY-SA
διαθέσιμο με άδεια CC-BY-NC-SA
διαθέσιμο με άδεια CC-BY-NC
Επιτρέπεται η επαναχρησιμοποίηση του έργου και η δημιουργία παραγώγων αυτού με απλή αναφορά του δημιουργού.
Επιτρέπεται η επαναχρησιμοποίηση του έργου με αναφορά του δημιουργού, και διάθεση του έργου ή του παράγωγου αυτού με την ίδια άδεια.
Επιτρέπεται η επαναχρησιμοποίηση του έργου με αναφορά του δημιουργού.
Δεν επιτρέπεται η εμπορική χρήση του έργου.
Επιτρέπεται η επαναχρησιμοποίηση του έργου με αναφορά του δημιουργού
και διάθεση του έργου ή του παράγωγου αυτού με την ίδια άδεια.
Δεν επιτρέπεται η εμπορική χρήση του έργου.
διαθέσιμο με άδεια CC-BY-ND
Επιτρέπεται η επαναχρησιμοποίηση του έργου με αναφορά του δημιουργού.
Δεν επιτρέπεται η δημιουργία παραγώγων του έργου.
διαθέσιμο με άδεια CC-BY-NC-ND
Επιτρέπεται η επαναχρησιμοποίηση του έργου με αναφορά του δημιουργού.
Δεν επιτρέπεται η εμπορική χρήση του έργου και η δημιουργία παραγώγων του.
διαθέσιμο με άδεια
CC0 Public Domain
διαθέσιμο ως κοινό κτήμα
Επιτρέπεται η επαναχρησιμοποίηση του έργου, η δημιουργία παραγώγων αυτού και η εμπορική του χρήση, χωρίς αναφορά του δημιουργού.
Επιτρέπεται η επαναχρησιμοποίηση του έργου, η δημιουργία παραγώγων αυτού και η εμπορική του χρήση, χωρίς αναφορά του δημιουργού.
χωρίς σήμανση
Συνήθως δεν επιτρέπεται η επαναχρησιμοποίηση του έργου.

## Slide 24: Διατήρηση Σημειωμάτων
Οποιαδήποτε αναπαραγωγή ή διασκευή του υλικού θα πρέπει να συμπεριλαμβάνει:
  το Σημείωμα Αναφοράς
  το Σημείωμα Αδειοδότησης
  τη δήλωση Διατήρησης Σημειωμάτων
  το Σημείωμα Χρήσης Έργων Τρίτων (εφόσον υπάρχει)
μαζί με τους συνοδευόμενους υπερσυνδέσμους.

## Slide 25: Χρηματοδότηση
Το παρόν εκπαιδευτικό υλικό έχει αναπτυχθεί στo πλαίσιo του εκπαιδευτικού έργου του διδάσκοντα.
Το έργο «Ανοικτά Ακαδημαϊκά Μαθήματα στο ΤΕΙ Αθήνας» έχει χρηματοδοτήσει μόνο την αναδιαμόρφωση του εκπαιδευτικού υλικού.
Το έργο υλοποιείται στο πλαίσιο του Επιχειρησιακού Προγράμματος «Εκπαίδευση και Δια Βίου Μάθηση» και συγχρηματοδοτείται από την Ευρωπαϊκή Ένωση (Ευρωπαϊκό Κοινωνικό Ταμείο) και από εθνικούς πόρους.